## Slide 1: SNF RRAM Standard Process Flow and Characterization of Lesker-2 TiN
Project Members: Akash Levy, Jung-Soo Ko
Staff Mentors: Michelle Rincon, Vijay Narasimhan, J Provine, Usha Raghuram
ENGR241 Week 8 Updates

## Slide 2: Aim & Goals
  Test and characterize ALD RRAM devices
  Develop TiN recipe on new Lesker with low oxygen contamination
  Compare TiN properties (composition, resistivity, topography)
  Make electrical/imaging data and test structure layouts available to SNF
  Make SOPs for RRAM stacks and report TiN findings
Process Flow Development
Parameter Optimization
Comprehensive Documentation
SNF Database

## Slide 3: Progress Updates: Fabrication
  Clean and grow SiO2 for isolation
  Sputter 30nm TiN (Sharpie for step)
    AJA (non-SNF)
    Lesker-1 (SNF)
    Lesker-2 (SNF)
  Characterization
    XPS
    Prometrix
    AFM
    Alphastep
Si
SiO2
TiN

## Slide 4: TiN Properties
 | AJA: A1 | Lesker-1: X1 | Lesker-1: X2 | Lesker-1: X3 | Lesker-1: X4 | Lesker-2: Y1
Target | Ti | TiN | TiN | TiN | TiN | TiN
Base Pressure | ~1e-7 Torr | ~1e-6 Torr | ~1e-6 Torr | ~1e-6 Torr (?) | ~1e-6 Torr | ~1e-7 Torr
RF Power | 200W | 150W | 200W | 200W | 200W | 200W
Gettering/Chamber | 2 min TiN ↑ | 2 min TiN ↑ | 5 min Ti pre, 5 min TiN ↑ | 10 min Ti pre, 5 min TiN ↑ | 5 min Ti pre, 5 min TiN ↑ | 5 min TiN pre, 5 min TiN ↑
Wafer Bias | 100V | 100V | 100V | 100V | 100V | 100V
Ar/N Flow | 30/3 sccm ↓ | 50/5 sccm ↓ | 30/3 sccm ↓ | 30/3 sccm ↓ | 30/0 sccm ↓ | 11.5/0 sccm ↑
Pressure | 2.15 mTorr | 35->5 mTorr | 30->3 mTorr | 20->2 mTorr | 20->2 mTorr | 15->2 mTorr
Sheet Resistance | 15.650 Ω/□ | FAILED | 2428.0 Ω/□ | 304.8 Ω/□ | 129.9 Ω/□ | 
Resistivity | 67.25 uΩ cm | FAILED | 6313 uΩ cm | 927 uΩ cm | 519.6 uΩ cm | 
Uniformity/Color | Excellent | Bad | Good | Good | Good | 
Thickness/Time | 43nm/1800s | ?nm/1800s | 26nm/900s | 30.4nm/1800s | 40nm/1800s | 
Roughness | 0.67nm | Won’t do | ? | ? | ? | 
Composition (Ti:N:O) | 46:51:3 | 6:4:90 | 38:39:23 | 43:43:14 | 43:43:14 | 

## Slide 5: RRAM Process Flow Outline
Si
Si
Pt
SiO2
Start with Si wafer
Standard RCA Clean
Deposit 10nm Ti, 30nm Pt (AJA),
150nm SiO2 (ccp-dep)
Si
Pt
SiO2
Pattern trenches (PT-Ox)
Post-metal clean
SiO2
Si
Pt
SiO2
SiO2
HfO2
Deposit 10nm HfO2 (Fiji)
Si
Pt
SiO2
SiO2
HfO2
TiN
Sputter 30nm TiN
Si
Pt
SiO2
SiO2
HfO2
TiN
Pattern top electrodes (PT-MTL)
Post-metal clean
Active device area
(Not drawn to scale)

## Slide 6: Progress Updates: New Samples
  RCA clean wafer
  Grow 50nm SiO2 with thermco4 (20 min. of 900C wet growth)
  Deposit 10nm Ti on AJA (0.5Å/s)
  Deposit 30nm Pt on AJA (1Å/s)
  Deposit 150nm SiO2 with ccp-dep (SiO350-0, SiOx1 350C, 2.5 min)
  Woollam to check SiO2 thickness
    Measure 55-75 degrees, use n-k fit for SiO2
    S1: 2320Å, R1: 1504Å
  Photolithography (1)
    YES Oven default recipe
    SVG Coat 2 with 1µm Shipley 3612 resist, 2mm EBR, skip vapor prime step
    Expose with Heidelberg (use defoc: -2, dose: 70/80 for Heidelberg1/2, 350nm laser src)
    SVG Develop (1µm 3612 dev) w/post-exposure bake and inspect
  PT-Ox: Ox_ER_Test, 210 sec.
  Resist ashing: Matrix default recipe
  Metal cleaning on wbflexcorr (1)
    SRS-100 for stripping photoresist from metal wafers, 60C, 15 min
    PRS-1000 for final metal clean before deposition, 40C, 10 min
  Alphastep to check etching
    Matched Woollam extremely well
  Fiji1 thermal HfO2: 50 cyc. condition,100 cyc. @ 200C
  Woollam to check HfO2 thickness
    Measure 55-75 degrees, use n-k for SiO2 from before thickness and n-k from before
    S1: 182Å, R1: 229Å, R2: 233Å, R3: 130Å
    Strangely high deposition rate
  TiN deposition 30nm
  Photolithography (2)
  PT-MTL: Al-HiPRSel-Usha, 30s
    Should etch through all of TiN and some oxide
  Resist ashing: Matrix default recipe
  Microscope to ensure TiN is patterned correctly
  Photolithography (3)
  PT-Ox: Ox_ER_Test, 210 sec.
  Resist ashing: Matrix default recipe
  Microscope to ensure Pt contact is patterned correctly
  Metal cleaning on wbflexcorr (2)
  Electrical testing
Notes: Dose: 70, defoc: -2 on Heidelberg-2
Dose: 80ish, defoc: -2 on Heidelberg-1

## Slide 7: Sample with AJA TiN: looks a bit weird

## Slide 8: Issues with Lesker-2
  Pressure rampdown causes deignition
  Upstream control issue
  Finding the sweet spot for deposition is tricky

## Slide 9: Budget
svgcoat | Jan-20 | 75 | 100
svgdev | Jan-20 | 30 | 40
wbflexcorr-1 | Jan-20 | 135 | 180
woollam | Jan-20 | 43 | 35.83
fiji2 | Feb-20 | 372 | 310
wbflexcorr-1 | Feb-20 | 42 | 35
wbflexcorr-3 | Feb-20 | 262 | 218.33
woollam | Feb-20 | 14 | 11.67
afm-asylum | Jan-20 | 76 | 25.33
aja-evap | Jan-20 | 39 | 22.75
heidelberg | Jan-20 | 120 | 160
heidelberg2 | Jan-20 | 15 | 20
Platinum (Pt) - aja-evap (nm) | Jan-20 | 30 | 17.7
gasonics | Jan 2020 | 25 | 20.83
PT-Ox | Jan 2020 | 36 | 30
svgcoat | Jan 2020 | 75 | 100
svgcoat2 | Jan 2020 | 36 | 30
svgdev | Jan 2020 | 81 | 67.5
svgdev | Jan 2020 | 30 | 40
Wafer - L-test (100 mm) | Jan 2020 | 4 | 68
wbclean-1 | Jan 2020 | 85 | 70.83
woollam | Jan 2020 | 35 | 29.17
yes | Jan 2020 | 27 | 22.5
fiji1 | Feb 2020 | 3 | 2.5
fiji2 | Feb 2020 | 178 | 148.33
matrix | Feb 2020 | 37 | 30.83
prometrix | Feb 2020 | 34 | 28.33
PT-Ox | Feb 2020 | 49 | 40.83
svgcoat2 | Feb 2020 | 28 | 23.33
svgdev | Feb 2020 | 39 | 32.5
Tweezer box - Small | Feb 2020 | 1 | 1
Tweezers - Teflon-coated | Feb 2020 | 1 | 48
wbclean-1 | Feb 2020 | 74 | 61.67
woollam | Feb 2020 | 198 | 165
yes | Feb 2020 | 33 | 27.5
aja-evap | Jan 2020 | 87 | 50.75
heidelberg | Jan 2020 | 58 | 33.83
heidelberg | Jan 2020 | 120 | 160
heidelberg2 | Jan 2020 | 157 | 91.58
heidelberg2 | Jan 2020 | 15 | 20
lesker-sputter | Jan 2020 | 160 | 93.33
aja-evap | Feb 2020 | 84 | 49
heidelberg2 | Feb 2020 | 54 | 31.5
lesker-sputter | Feb 2020 | 260 | 151.67
lesker-sputter2 | Feb 2020 | 144 | -2.4
Platinum (Pt) - aja-evap (nm) | Feb 2020 | 30 | 17.7
XPS: PHI VersaProbe 3 | Jan 2020 | 152 | 126.67
XPS: PHI VersaProbe 3 | Jan 2020 | 165 | 220
XPS: PHI VersaProbe 3 | Feb 2020 | 223 | 130.08
 |  |  | 3438.97

## Slide 10: Thank you!
Questions?

## Slide 11: AFM Roughness Measurements: AJA
Roughness: 1.02nm SiO2
Roughness: 1.22nm TiN
0.67nm independent

## Slide 12: AJA TiN Sputter Recipe: Reactive Sputtering
  Target: Ti
  RF Power: 200W
    150W also ok; different dep rates
  Wafer Bias: 100V
    DC bias 92-94V on AJA in practice
  Ar Flow: 30 sccm
  N Flow: 3 sccm
    Unit conversion means that this is written as 30 on AJA somehow
  Set Pressure: 2.15 mTorr
    1.8-1.9 mTorr on AJA in practice
  Recipe:
  Load sample and wait for 1e-7 base chamber pressure
  Plasma on another target (conductive) e.g. Al, GeTe
    Eases plasma ignition for TiN target
  Plasma on Ti target
    Bias on other target off
  20s ramp-up of Ti target RF power
  2 minutes pre-sputtering of Ti target (shutter closed, no wafer bias)
  20s ramp-up of wafer bias
  Open shutter and deposit for X seconds
  Close shutter
  20s ramp-down of Ti target RF power

## Slide 13: Literature Review on TiN Sputtering
  TiN films absorb a high concentration of contaminants including hydrogen, carbon, and oxygen when they are exposed to air after deposition
  With the target–substrate distance set to 88 mm the contaminant levels increase from ∼0.1% to ∼10% as the pressure is increased from 2 to 9 mTorr
  The contaminant concentrations also correlate with in-plane distance from the center of the substrate and increase by roughly two orders of magnitude as the target–substrate distance is increased from 88 to 266 mm
  Contaminants strongly influence properties of TiN thin films
  Resistivity of stoichiometric films increases by around a factor of 5 as the oxygen content increases from 0.1% to 11%
  Energy of the sputtered TiN particles plays crucial role in determining the film properties
  Important to precisely control the energy of these particles to obtain high-quality TiN
Ohya, S., B. Chiaro, A. Megrant, C. Neill, R. Barends, Y. Chen, J. Kelly et al. "Room temperature deposition of sputtered TiN films for superconducting coplanar waveguide resonators." Superconductor Science and Technology 27, no. 1 (2013): 015009.